Zeiss LSM 710 User Guide

STARTUP

- Turn on Exfo lamp to the left of the microscope 1

- Turn on [Main] 2, [System/PC] 3 and [Component] 4 switches

- Turn on Argon laser (458, 488 or 514nm):

- Check that the toggle switch 6 is set to [idle]

- Turn the key on the front of the power supply clockwise 5

- After 5 minutes move laser toggle switch 6 to [run]. After another few minutes adjust the potentiometer knob 7 to the point just below where the red LED becomes illuminated. (While waiting, continue the startup sequence)

- Turn on computer 8 and wait for log on window

- On Exfo lamp 1 push on the Open button 9 once hour display stopped blinking

FOR THE FIRST USER OF THE DAY

log on using:		your HawkID

and password:	HawkID password

- Start [ZEN] software and do not use the microscope while the software loads.

VISUAL OPERATION

- Load slide

- Select a X5 or X10 objective lens. Make sure the stage is sufficiently lowered [Loading position] using the screen pad on the right of the microscope or use the lower button at the base of the instrument. Select [Set work position] on the screen pad.

-Load sample onto stage and focus on your sample then select [Set work position] icon on the screen pad.

Note: Do not depend on the preset load and work settings. Samples and slides of different thicknesses may have been used by the previous investigator.

- In the [Locate] tab select the desired fluorophore or DIC filter. Automatic shutter opening/closure upon selection of an imaging mode and TL shutter closes when reflected mode selected.

- Select the proper objective lens

- Locate region of interest

- Close the reflected or transmitted light shutter used. To turn off the TL lamp, click on the TL icon and then click on [ON] (ON in blue = lamp ON, TL in gray = TL lamp OFF).

CONFOCAL OPERATION

- Select [Acquisition]

- If you have a previously acquired image saved in the format of acquisition (.czi), open that image in the program and click on the [reuse] button at the bottom of the image window, or

C- Select [Smart Setup] and [Current] to keep full control of the confocal acquisition setup and select [Smartest Line], or [reuse] parameters from a previous .czi image file

D- The lasers, with the exception of the argon laser, will automatically turn ON if needed from the configuration setup or modify the excitation line under the [Laser] menu

ATTENTION

If you are using a method from the “Smart setup” requiring more than ONE track (as seen in the [Imaging setup] window, and not using the [Smartest Line] you must follow these guidelines:

- In the [Imaging setup] window select Switch track every [line] instead of [frame] using the pulldown window.

- Check that each track displays the proper highlighted emission curve. If not, adjust using the cursor under the emission spectra.

All the tracks selected must have the same setup in their MBS (seen in the [Light Path] window).

- Select [Expand all] under the [Channel] window. For each track loaded you will see the laser line necessary for each dye excitation.

- Under the [Light Path] window, select the MBS matching the excitation lines required for your imaging by clicking on the MBS icons (1st one and 2nd one under the window).

D- In the Acquisition Mode menu select the desired parameters for final image

E- Select [Auto Exposure] to get initial gain and offset settings

OR

F- Select [Live] to find your plan of focus using the focus knobs

G- In the [Channels] menu adjust the [laser power], [pinhole], [gain] and [offset] as needed

H- Select [Snap] to collect the final image

I- Save image to desired location (Do not store on drive C:\ instead use the D:\ drive)

J- Select .czi format.

ATTENTION

All images left on the D-drive are periodically purged (every 3 months) without warning

SHUTDOWN afterhours:

Check the online scheduler to determine if someone is scheduled after you (within 0.5 hour).

IF YES:

- Turn the Argon laser toggle switch to [idle], close your images BUT let program opened

IF NO:

- Turn the Argon laser toggle switch 6 to [idle] and turn off key 5

- Wait for Argon laser fan to switch off (about 5 min from turning the key off)

- Move stage to load position

- Remove slide

- Gently remove oil from oil lens using lens paper or cotton swabs. Use a cotton swab with a 		little isopropyl alcohol to remove final traces of oil

- Exit the Zen software then shut down the computer 8

- Turn off the [System] 3, [Components] 4 and [Main] 2 switches in this order

- Turn off the Exfo lamp 1

- Log your usage in the log book by the microscope.

The NIH grant number for this instrument should be included in all publications. Grant Number Zeiss LSM710   1 S10 RR025439-01